## Slide 1: Embedded System for Control and Data Acquisition of NLO Experiments
Presented by
Syed Razwanul Haque(Nabil)
Reg. no: 2011122011
Session: 2011-12
Supervised by
Prof. Yasmeen Haque
Department of Physics, SUST
Department of Physics, Shahjalal University of Science and Technology
Sylhet-3114, Bangladesh

## Slide 2
Two devices have been developed to control the laser and rotational stage.
AVR Microcontroller has been used where the programming language was C++.
The remote control has been developed for the laser controller and external mechanical device has been designed to mount a stepper motor with the rotational stage
The measured Accuracy for the controller of laser and rotational stage has been found to be 0.004% and 9% respectively.
Open Source Arduino Platform has been used
Abstract

## Slide 3
Outline of the Presentation
Title | Slide no.
The Embedded Systems | 1
Embedded System Architecture | 2
Microcontroller & Arduino | 3
Advantages of Embedded System | 4
Laser Controller | 5-14
Rotational Stage | 15-25

## Slide 4: Embedded Systems
An embedded system is a microcontroller/microprocessor based system that is built to control a function or a range of functions.
1
Embedded Systems talk with the outside world via peripherals, such as:
>Serial Communication Interfaces (SCI)
>Synchronous Serial Communication Interface
>Universal Serial Bus (USB)
>Multi Media Cards
> Field buses: CAN-Bus, LIN-Bus,
>Timers: Capture/Compare and Time Processing Units
>Discrete IO: aka General Purpose Input/output (GPIO)
>Analog to Digital/Digital to Analog (ADC/DAC)

## Slide 5
M1
Embedded System Architecture
2

## Slide 6
Microcontroller & Arduino
Micro
controller
Arduino
Arduino is an open-source electronics prototyping platform based on flexible, easy-to-use hardware and software. It's intended for artists, designers, hobbyists and anyone interested in creating interactive objects
or environments.
3
A microcontroller is a small computer on a single integrated circuit containing a processor core, memory, and programmable input/output peripherals

## Slide 7
Advantages of Embedded System
1) No external memory is required as the memory is embedded with the hardware.
2) It gives the enhanced speed because there is no switching as memory is embedded in it.
3) Due to less or no switching, power dissipation is also less.
4) Traditional Digital Circuits are not Programmable but MCU is Programmable
5) They allow the system hardware to be simplified so costs are reduced
6) Artificial Intelligence can be implemented in the Device
7) It is Standalone and circuit size is small
4

## Slide 8
Laser Controller
Problem With Existing Laser Controller:
Analog Control Method
Huge Time Consumption during Experiment
Laser Power Controlling Knob is not Precise
Lack of Safety
External Illumination Source is Required
Laser Power Increaser Exponentially With Knob
Touching the laser physically is not convenient
User cannot give numerical inputs
There is no display to show its operating data
Analog controller is not configurable
5

## Slide 9
How Embedded System Makes it Better
6
Digital Control Mechanism
Time Consumption Reduces by 50 Times than Previous One
Wireless Remote Control System
On Board Operating System
Menu System for Interactive User Interface
Password Locking System
Laser Intensity Can be Controlled Precisely
LCD Display Interface
SD Card for Data Saving
Real Time Clock Interface
Protection Algorithm for Safe Operation
Increase or decrease power linearly
Ports for installs different kinds of modules
External Light Source is Not Required

## Slide 10
Developed Laser Controller
Photograph of Laser Remote Unit
7
Photograph of Laser Control Unit

## Slide 11: System Architecture of Laser Controller
Fig: Laser Control Unit
Fig: Remote Control Unit
8
User Input
&
Display
Wireless
Transceiver
Command
Processing
Laser
Laser
Control
Circuit
Command
Processing
Wireless
Transceiver

## Slide 12
Menu System
9
Power Must Be Zero
Remote Unit
Remote ON
(When ON/OFF switch pressed)
(Initial Screen)
Enter Password
0----
(4 digit password)
Wrong Password
Unlocked
(If password correct)
Press OK For Menu
Laser Control
Settings Back
A1
(If OK pressed)
(If password wrong)
[Laser Control]
[SH-OFF] POW-0
LAS_1_OFF LOCK
Back
Are You Sure Yes No
Laser 1 Running
!! LOCKED !!
Laser OFF
(If Power Zero)
(If Power>0)
A3
A2

## Slide 13
Fig: Laser Remote Unit Circuit
Laser Controller Circuit
10

## Slide 14
For λ= 457nm
Experiments Using Laser Controller
11
Controller Indicator (%) | Laser Power(mw) | Std. dev(µw )
0 | 0.368 | 2.9784
10 | 0.630 | 0.3265
20 | 1.086 | 1.8903
30 | 1.533 | 1.4437
40 | 1.982 | 2.2236
50 | 2.425 | 1.1493
60 | 2.868 | 1.0683
70 | 3.317 | 1.4082
80 | 3.761 | 1.9118
90 | 4.208 | 2.2376
100 | 4.643 | 2.0995

## Slide 15: The Graphs (Laser Controller)
12

## Slide 16
Conclusion (Laser Controller)
The developed device can be controlled remotely using radio signal thus reducing human touch related error
Time consumption for changing intensity manually was approximately fifty times greater than the newly implemented remote control system
The using remote control intensity controller change in intensity can be controlled precisely unlike the manual process
Error margin for the measurement using laser controller is 0.004%
13

## Slide 17: Future Works (Laser Controller)
14
Xbee module for more reliable wireless communication.
Wi-Fi module for smart phone interfacing.
Touch screen display to show graph in remote control unit.
Ethernet communication to interface laser with central network.
GSM/GPRS module to monitor and transfer data from a remote distance.
Digital charging system
PC interfacing
Power meter interfacing with laser controller
Rotational stage controller interfacing with laser controller
Temperature sensor, humidity sensor, gas sensor for more safety

## Slide 18: Rotational Stage
15
The rotational stage is a mechanical device which can hold and rotate an optical component to perform many optical experiments
The optical element is secured inside the central aperture of the platform of the holder by the threaded retaining ring
This enables the user to set the current position on the scale to any chosen angle independently of the position of the holder
The position is marked in degree from 0 to 360 degrees
Angular position is read on the 360° graduated scale on the platform, which is marked every 2°
It can be rotated either manually or by means of a motorized system
In this Research work a microcontroller based system converts a manual rotational stage into a motorized rotational stage

## Slide 19: Problem with Manual Rotational Stage
16
A manual rotational stage has no control unit and it is also not motorized
Rotating this stage by hand is time consuming in addition to being a waste time during experimentation
External Light Source is Required to Operate
If data is taken for every 2° rotation the knob needs to be touched and the flash light switched on and off 180 times
Digital Input is Not Possible
Wireless Control is Not Possible
Automation is not possible
Interfacing with External Device is not Possible
Manual Rotational Stage Consumes at least 3 Hours for a single experiment if resolution is 2 degree

## Slide 20
17
How Embedded System Makes it Better
Digital Control Mechanism
Onboard Operating System
Automatic Error Correction
Interfacing with External Device
External Light Source is Not Required
Quick Operation
Remote Control
Stable Operation
User Friendly
Digital Calibration System

## Slide 21
Block Diagram of Developed Rotational Stage
18

## Slide 22
Photograph after Motorized
19
Fig: Mechanical Unit
Fig: Remote Unit

## Slide 23
Engineering Drawing
Fig: Before Motorized
Fig: After Motorized
20

## Slide 24
Circuit Diagram
Fig: MOSFET H-Bridge
Fig: Control Unit Circuit
21

## Slide 25
Experimental Data
22
Angle of Rotation | Actual Rotation | Error | Error Percentage
(in degrees) | (in degrees) | (in degrees) | %
360 | 393.00 | 33.00 | 9.17%
 | 393.00 | 33.00 | 9.17%
 | 390.00 | 30.00 | 8.33%
 | 393.00 | 33.00 | 9.17%
 | 391.00 | 31.00 | 8.61%
270 | 294.00 | 24.00 | 8.89%
 | 293.00 | 23.00 | 8.52%
 | 294.00 | 24.00 | 8.89%
 | 294.00 | 24.00 | 8.89%
 | 294.00 | 24.00 | 8.89%
180 | 195.00 | 15.00 | 8.33%
 | 197.00 | 17.00 | 9.44%
 | 195.00 | 15.00 | 8.33%
 | 197.00 | 17.00 | 9.44%
 | 195.00 | 15.00 | 8.33%
120 | 131.00 | 11.00 | 9.17%
 | 132.00 | 12.00 | 10.00%
 | 131.00 | 11.00 | 9.17%
 | 131.00 | 11.00 | 9.17%
 | 131.00 | 11.00 | 9.17%
90 | 98.00 | 8.00 | 8.89%
 | 97.00 | 7.00 | 7.78%
 | 98.00 | 8.00 | 8.89%
 | 98.00 | 8.00 | 8.89%
 | 98.00 | 8.00 | 8.89%

## Slide 26
Error Graph
23

## Slide 27
Conclusion (Rotational Stage)
24
  Changing rotation manually for rotational stage needs extra light source and also needs human touch which is time consuming and also includes measurement error. Developed digitally controlled rotational stage saves significant time and makes experimentation more precise.
Previous Time Consumption
Current Time Consumption
180 minutes
5 minutes

## Slide 28
Future Works (Rotational Stage)
Wireless control feature
Artificial Intelligence
Graphical LCD to show more information during operation
Automatic error reduction system
Computer interfacing option
Internet based control system
Interfacing with Laser Controller Device
Smartphone Interfacing
25

## Slide 29
Thanks
26